## Cover Sheet
Top Level Eng Specification |  | CMD PN M.C00.00.C, Rev 0 | [5 empty] | Inspection Record # | (XXXX) PO # Lot # In process Insp
Cage Aprevo ACDF-X
Part # |  | {{Item 1}} | {{Item 10}} | Customer Lot # | {{Lot Number}} |  |  | Total Part Qty | {{Total Quantity}}
|  | {{Item 2}} | {{Item 11}}
|  | {{Item 3}} | {{Item 12}}
|  | {{Item 4}} | {{Item 13}}
|  | {{Item 5}} | {{Item 14}}
|  | {{Item 6}} | {{Item 15}}
|  | {{Item7}} | {{Item 16}}
|  | {{Item 8}} | {{Item 17}}
|  | {{Item 9}} | {{Item 18}}
Master Job # |  | {{Job Number}} |  | Traveler # |  |  |  | Date Inspected
Notes
ID | Inspection Criteria | Inspection Method or Tool | Nominal Spec Value & Tolerance | Drawing Sheet | Zone / Note | AQL Level or 100% | Inspection Qty | Rejected Qty | Notes
1 | Perform line clearance | Visual | Pass / Fail | N/A | N/A | 100% |  |  | N/A
2 | Verify lot count | Visual | Pass / Fail | N/A | N/A | 100% |  |  | N/A
3 | Verify parts are free from visible damage, defects, cracks, & burrs | Visual | Pass / Fail | N/A | N/A | 100% |  |  | N/A
4 | Verify parts are clean (free from debris & contamination) | Visual | Pass / Fail | N/A | N/A | 100% |  |  | N/A
5 | Surface Finish: Glass Bead Blast | Visual | Pass / Fail | 1 | Note 4 | 100% |  |  | N/A
6 | All support peaks are <0.25mm | Visual | Pass / Fail | 1 | Note 10 | 100% |  |  | N/A
7 | Patient Side "R": Present and Correct | Visual | Pass / Fail | 1 | Note 15 | 100% |  |  | N/A
8 | Markings on the parts "XXXXX" for the lot number, "CXX"for implant level, "Plan" doesn't include a suffix, "-" is the suffix for "Minus", are present and correct | Visual | Pass / Fail | 1 | Note 15, C 1-3, C 1-1 | 100% |  |  | N/A
9 | Dimension A | Calipers | See Form 19 from the customer ______ ± 0.5 | 1 | C 1-2 | 100% |  |  | *Record the sample measurements on Measurement Sheet
10 | Dimension B | Calipers | See Form 19 from the customer ______ ± 0.5 | 1 | D 1-2 | 100% |  |  | *Record the sample measurements on Measurement Sheet
11 | Dimension C | Calipers/Keyence | See Form 19 from the customer ______ ± 0.5 | 1 | C 1-3 | 100% |  |  | *Record the sample measurements on Measurement Sheet
12 | Threads formed properly | Thread Gage M2.0X0.40 | Pass / Fail | 1, 3 | B 1-2, C 3-1 | 100% |  |  | N/A
13 | 2x Ø3.73 mm | Pin Gage | 3.73± 0.10 | 1 | C 1-3 | 100% |  |  | *Record measurements on Measurement Sheet
14 | 2x Ø4.35 mm | Pin Gage | 4.35± 0.10 | 1 | C 1-3 | 100% |  |  | *Record measurements on Measurement Sheet
15 | Dimension 1.20mm | Drop gage | 1.20± 0.10 | 3 | C 3-1 | 100% |  |  | *Record measurements on Measurement Sheet
16 | Ø2.81 mm | Calipers | 2.81± 0.10 | 3 | C 3-3 | 100% |  |  | *Record measurements on Measurement Sheet
17 | Ø2.10 mm | Pin Gage | 2.10± 0.10 | 3 | C 3-3 | 100% |  |  | *Record measurements on Measurement Sheet
18 | Dimension 1.25mm | Calipers | 1.25± 0.10 | 3 | C 3-2 | 100% |  |  | *Record measurements on Measurement Sheet
19 | Dimension 1.75mm | Calipers | 1.75± 0.10 | 3 | C 3-2 | 100% |  |  | *Record measurements on Measurement Sheet
| Inspection Tool | Equipment # | Cal Due Date |  | Lot Size | < 8 | 8 to 15 | 16 to 25 | 26 to 50 | 51 to 90 | Over 500
Calipers
| Keyence |  |  |  | 1.0 AQL Sample Size* | 100% | 13 | 13 | 13 | 13 | Contact QA
Pin Gage
Drop gage
*All AQL sampled parts must meet acceptance criteria or 100% inspection is required.
_________________________________________________________________
Inspected by:
Form Revision History (F-825-1052)
| Rev | DCO # | Effective Date | Description of Change
| C | 25-083 | 15Jul2025 | Update ID 15 Dim value from 1.0 to 1.2, added additional columns in measurement sheet for ID 13 & 14.
| B | 25-076 | 08Jul2024 | Update the title of the form, include Keyence measurement option for dimension C
| A | 25-071 | 30June2025 | Initial release
Template Revision History (F-825-1000)
| Rev | DCO # | Effective Date | Description of Change
| C | 25-051 | 21Apr2025 | Updated doc title, generally updated field sizing/locations, added SpiTrex 3D logo, added Inspection Record field, clarified instructional notes, added ID column, added "or 100%" to AQL heading, added Measurement Sheet, added rev history for forms, moved template number to footer
| B | 25-033 | 19Mar2025 | Removal of the second signature, removal of the process control tolerance
| A | 23-054 | 23Jul2023 | Initial Release with new numbering system (F-825-1XXX and F-825-2XXX)

## Measurement Sheet
ID | 9 | 10 | 11 | 13 | 13 | 14 | 14 | 15 | 16 | 17 | 18 | 19
Inspection Criteria | Dimension A | Dimension B | Dimension C | Ø3.73 mm | Ø3.73 mm | Ø4.35 mm | Ø4.35 mm | Dimension 1.20mm | Ø2.81 mm | Ø2.10 mm | Dimension 1.25mm | Dimension 1.75mm
Nominal Spec Value & Tolerance | See Form 19 from the customer ______ ± 0.5 | See Form 19 from the customer ______ ± 0.5 | See Form 19 from the customer ______ ± 0.5 | 3.73± 0.10 | 3.73± 0.10 | 4.35± 0.10 | 4.35± 0.10 | 1.20± 0.10 | 2.81± 0.10 | 2.10± 0.10 | 1.25± 0.10 | 1.75± 0.10
Sample
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35